## Лист1
Done | Tested | Tested on hardware | Comm | send | receive | Function
|  |  | SC: | xx |  | set mA setpoint value |  |  | set ma | ma
|  |  | CN |  | xxxxxxxxxx | read mA setpoint value |  |  | read kvma
|  |  | CA |  | xxxxxxxxxx | read mA actual value
|  |  | SV: | xx |  | set kV setpoint value |  |  | set kv | kv
|  |  | VN |  | xxxxxxxxxx | read kv setpoint value |  |  | read kvma
|  |  | VA |  | xxxxxxxxxx | read kV actual value
|  |  | SN: | xx,xx |  | set kV and mA setpoint value |  |  | set kvma | kv,ma
|  |  | GN |  | xxxxxx:xxxxxx | read kV and mA setpoint |  |  | read kvma
|  |  | GA |  | xxxxxx:xxxxxx | read kV and mA actual value
|  |  | TS: | x |  | exposure timer ON | x = number of exposure timer |  | timer on | number
|  |  | TE: | x |  | exposure timer OFF | x = number of exposure timer |  | timer off | number
|  |  | TP: | x,hh,mm,ss |  | set exposure time setpoint value | x = number of exposure timer | hh - hours etc. | set timer number time | number,hh,mm,ss
|  |  | TN: | x | yyyyyyyyyy | read exposure timer setpoint value | x = number of exposure timer | y - seconds | read timer s | number
|  |  | TA: | x | yyyyyyyyyy | read exposure timer actual value | x = number of exposure timer | y - seconds |  | number
|  |  | HV: | x |  | high voltage ON/OFF | x = 0 - OFF (default), x = 1 - ON |  | voltage | on/off
|  |  | SR: | xx | yyyyyyyyyy | read status word | xx - number of status word | y - return | read status
|  |  | SW: | xx,yyy |  | write status word | xx - number of status word | yyy - data of status word (DEC) | set status
|  |  | ER: | xx | ascii | read current message | xx - number as per message table |  | read msg
|  |  | CL |  |  | acknowledge error
|  |  | WU: | x,yy |  | warm up programm | x = non-operative interval x = 0 - no warm up x = 1 - 1day x = 2 - 2day x = 3 - 1week x = 4 - warm up via RTC yy = test voltage kV |  | warm up | number,test voltage
|  |  | SP: | xxxx |  | set tube output Watts |  |  | set power | watts
|  |  | RP |  | xxxxxxxxxx | read tube output Watts |  |  | read power
|  |  | KB: | x |  | input via control module | x = 0 - input disable x = 1 - input enable (default) |  | input module | enable/disable
|  |  | SF: | x |  | select focus | x = 1 - 0.15 x 8 mm x = 2 - 0.4 x 8 mm x = 3 - 0.4 x 12 mm x = 4 - 1.0 x 10 mm x = 5 - 2 x 12 mm x = 6 - 0.04 x 8 mm x = 7 - 0.04 x 12 mm x = 8 - 0.1 x 10 mm x = 9 - 0.2 x 12 mm x = 10 - 0.4 - 0.08 mm x = 11 - 0.4 x 1.2 mm x = 12 - 1 x 1 mm x = 13 - 2 x 1.2 mm |  | set focus | 1-13
|  |  | RF |  | xxxxxxxxxx | focus |  |  | read focus
|  |  | FR |  |  | read focus | [string]
|  |  | SM: | x |  | select anode material | x = 1 - Ti x = 2 - Cr x = 3 - Fe x = 4 - Co x = 5 - Cu x = 6 - Mo x = 7 - Ag x = 8 - W x = 9 - Au |  | set material | ti - au
|  |  | RM |  | xxxxxxxxxx | read anode material |  |  | read material
|  |  | MR |  |  | read anode material | [string]
|  |  | RH |  | xxxxxxxxyy | elapsed operating hours of equipment elapsed operating hours of current tube | xxxxxxxx - hours, y - 0.01 of hour |  | read tube
|  |  | OS: | x |  | open beam shutter |  |  | shutter open | number
|  |  | CS: | x |  | close beam shutter |  |  | shutter close | number
|  |  | CC: | wxyz |  | Sbeam shutter control | w = 1shutter, x = 2sh, y = 3sh, z = 4sh ; 0 - control via hands 1 - control via computer |  | shutter control | 1-4, 0-1
|  |  | HO: | xxx | dsetse | operating history |  |  | history
|  |  | HW: | xxx | dsetse
|  |  | ID: |  |  | read software id |  |  | id
|  |  | XC: |  | xxxxxxxxxx | read tube number |  |  | read tube
|  |  | WT: |  | xxxxxxxxxx | warm up time left
|  |  | CT: | x |  | tube change | x = tubes 1-8 |  | tube change | number
|  |  | LS: | x |  | language selection | x = 1 hitler, x = 2 winston |  | set language | de,en,fr,es
|  |  | DR |  | 0000uvwxyz | read date | u = day, v = month, w = year, x - h, y - m, z - s |  | date
|  |  | DS |  | dd,mm,yy;hh,mm,ss | set current date |  |  | set date | dd,mm,yy,hh,mm,ss